MRI Prostate Segmentation Assistant (MRI PSA)

Technical text and diagrams for the SIG Final Report

Prepared against “SIG Final Report – DRAFT v0.9”

How to use this document: The pale-blue boxes contain final prose intended for direct copying into the SIG report. The yellow boxes are author-only cautions and should not be pasted as project results. The source draft was not modified.

Scope decision: The CARPL screenshot and Figure 3 caption are deliberately deferred because MRI PSA has not yet been viewed end-to-end on the CARPL platform.

# 1. Copy-ready replacement text

## A. Replace the highlighted architecture placeholder

Model development. After radiologist review, the whole-gland and lesion segmentation models were trained on CHROMA using the curated SGH cohort and the public PROSTATEx dataset (Table 1). The whole-gland model is a three-dimensional U-Net implemented in MONAI. It receives axial T2-weighted MRI as a single input channel and produces a two-class background/prostate probability map. The lesion model is an ensemble of five three-dimensional residual RR-UNets. Each fold receives co-registered T2-weighted, apparent diffusion coefficient (ADC) and highest-b-value diffusion images and produces a two-class background/lesion probability map. During inference, T2-weighted images are resampled to 1.0-mm isotropic spacing for whole-gland segmentation, while the three lesion channels are aligned and resampled to 0.5-mm isotropic spacing. Lesion inference is restricted to a prostate-centred region of interest. The five fold probabilities are averaged, constrained by the cleaned whole-gland mask and thresholded at 0.63 to generate the final lesion mask. A separate three-dimensional ResNet assigns an AI PI-RADS category from 2 to 5 to each detected lesion component. Training was completed on 31 August 2026. The planned deployment uses the binary whole-gland model; the separately trained transition-zone/peripheral-zone model is not included.

## B. Replace the highlighted results sentence

The whole-gland model exceeded the prespecified 90% Dice target. On the 264-case development-validation cohort, it achieved a Dice score of 0.95447, precision of 0.95228 and recall of 0.95781. For lesion segmentation, five-fold out-of-fold evaluation across 1,189 unique cases produced a Dice score of 0.47891, lesion sensitivity of 73.62% (773 of 1,050 reference lesions detected) and 0.743 false-positive lesions per case. These are development-validation results used to characterise model performance; they are not results from the separate locked test cohorts.

Locked-test boundary: The project documents define separate locked test cohorts of 264 whole-gland cases and 209 lesion cases. No completed locked-test metrics were found in the reviewed report, approval form, repository evaluation outputs or training artefacts. Do not relabel the 0.95447 and 73.62% results as locked-test performance.

## C. Optional sentence immediately after the results paragraph

The predefined locked test cohorts comprised 264 whole-gland cases and 209 lesion cases and were kept separate from model selection. Final locked-test evaluation remains to be completed or documented before those results are reported.

## D. Replace the first two sentences under “Integration and deployment”

The inference workflow is packaged as a Dockerised MONAI Deploy application for planned onboarding to CARPL. It generates a cleaned whole-gland mask, lesion mask, per-lesion AI PI-RADS output, DICOM Structured Report and a combined DICOM RT Structure Set referenced to the selected T2-weighted series. The combined RTSTRUCT has been tested with Biobot’s updated Urofusion and was successfully loaded with separate prostate and lesion structures. At the reporting cut-off, end-to-end CARPL deployment and prospective clinical validation had not yet been completed and remained dependent on AI Office approval and the relevant platform and cyber-security clearances.

Why this wording matters: The draft currently states that the model “runs on CARPL”. The repository and the draft’s own problem/milestone sections instead support “packaged/prepared for CARPL onboarding”; live end-to-end CARPL use remains pending.

# 2. Technical model specification

| Component | Verified implementation | Inference processing | Output |
| --- | --- | --- | --- |
| Whole-gland segmentation | 3D MONAI U-Net; 1 T2 input channel; 2 output channels (background/prostate); feature channels 16–32–64–128–256; four downsampling stages; two residual units per stage. | RAS orientation; 1.0-mm isotropic resampling; channel-wise z-score normalisation; 128×128×16 sliding window, 50% overlap, batch size 4; softmax and argmax in original image space. | Binary whole-gland mask; disconnected islands removed and internal holes filled for downstream use. |
| Lesion segmentation | Five-model ensemble of 3D residual RR-UNets; 3 inputs (T2, ADC, highest-b DWI); 2 outputs (background/lesion); 32 initial filters; encoder blocks 1–2–3–4; decoder blocks 3–2–1; residual connections; no recurrent, attention or squeeze-and-excitation blocks in the deployed configuration. | ADC and high-b volumes aligned to T2; all channels resampled to 0.5 mm; prostate-centred crop with 32-voxel margin; per-channel z-score normalisation; five probability maps averaged; cleaned-gland constraint; threshold 0.63. | Binary lesion mask plus five fold probability maps and a merged probability map. |
| Lesion classification | 3D ResNet with BasicBlock stages [1,1,1,1], channels 32–64–128–256 and four output classes. This classifier is downstream of segmentation and was not part of the segmentation performance table. | 64×64×64 crop around each 3D lesion component; class index mapped to PI-RADS 2–5; a documented size rule can promote PI-RADS 4 to 5. | Per-lesion AI PI-RADS category, lesion volume and major-axis length. |
| Clinical-format export | Canonical combined DICOM RTSTRUCT plus DICOM TID 1500 measurement SR, generated from the masks and selected source images. | Every contour is referenced to the actual selected T2 SOP instance and checked for source membership, physical plane agreement and mask round trip before publication. | Prostate ROI first, followed by separate lesion ROIs; AI-generated output remains UNAPPROVED/UNVERIFIED pending radiologist review. |

# 3. Copyable architecture diagram

Figure X. MRI PSA inference workflow

| DICOM study T2W + ADC + DWI | → | Series selection + highest-b filtering |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| WHOLE-GLAND BRANCH |  | T2W 1.0-mm isotropic | → | 3D MONAI U-Net sliding window | → | Cleaned whole-gland mask |
| LESION BRANCH |  | T2W + ADC + high-b 0.5-mm; gland ROI | → | 5 × 3D residual RR-UNet folds | → | Mean probability × gland; threshold 0.63 |
| CHARACTERISATION |  |  |  | 3D ResNet lesion crop | → | AI PI-RADS 2–5 per lesion |
| Masks | → | Measurements | → | DICOM RTSTRUCT + DICOM SR | → | Radiologist review in Urofusion |

Figure X. MRI PSA inference workflow. T2-weighted MRI drives whole-gland segmentation by a 3D U-Net. Co-registered T2-weighted, ADC and highest-b-value diffusion images, together with the cleaned whole-gland mask, drive a five-fold 3D residual RR-UNet lesion ensemble. The fold probabilities are averaged, constrained to the gland and thresholded to form the lesion mask. A downstream 3D ResNet assigns an AI PI-RADS category to each detected lesion. The masks and measurements are exported as DICOM RTSTRUCT and DICOM SR for radiologist review.

# 4. Results and validation boundary

| Model | Evaluation set actually reported | Dice | Additional reported metrics | Safe label |
| --- | --- | --- | --- | --- |
| Whole gland | 264 development-validation cases | 0.95447 | Precision 0.95228; recall 0.95781 | Development-validation |
| Lesion | Five-fold out-of-fold evaluation; 1,189 unique cases | 0.47891 | Sensitivity 73.62% (773/1,050 lesions); 0.743 false positives/case | Out-of-fold development validation |
| Whole gland — locked test | 264 reserved cases | Not located | Not located | Do not report until verified |
| Lesion — locked test | 209 reserved cases | Not located | Sensitivity and FP/case not located | Do not report until verified |

### Locked-test insertion template (complete only from the evaluation summary)

On the locked test cohorts, the whole-gland model achieved a Dice score of [INSERT VERIFIED VALUE] across 264 cases. Across the 209-case lesion locked test cohort, the ensemble achieved lesion sensitivity of [INSERT VERIFIED VALUE] ([DETECTED]/[REFERENCE] lesions), [INSERT VERIFIED VALUE] false-positive lesions per case and a Dice score of [INSERT VERIFIED VALUE]. The locked test cohorts were not used for model or threshold selection.

Figure Y. Evaluation boundary

| Model training SGH + PROSTATEx | → | Validation / OOF reported metrics | → | Locked test reserved; metrics not located | → | Prospective CARPL trial not yet completed |
| --- | --- | --- | --- | --- | --- | --- |

# 5. Figure 3 / CARPL screenshot

KIV as requested: Do not insert an MRI PSA-on-CARPL screenshot or caption yet. A valid figure should be captured only after an actual end-to-end CARPL run has received a study, executed inference, returned the RTSTRUCT/SR and displayed the result against the correct study.

## Caption to use later, once directly evidenced

Figure 3. MRI PSA output displayed after end-to-end processing on the CARPL platform. The combined DICOM RTSTRUCT contains the whole-gland contour and separate lesion structures; the radiologist reviews and edits these structures in Urofusion before biopsy planning.

# Appendix — Evidence map for final verification

| Claim | Primary evidence reviewed |
| --- | --- |
| Whole-gland architecture and inference | configs/organ.yaml; training/engine.py; prostate_mri_lesion_seg_app/organ_seg_operator.py; deployed models/organ/model.ts parameter shapes. |
| Lesion architecture and inference | configs/lesion.yaml; prostate_mri_lesion_seg_app/custom_lesion_seg_operator.py; prostate_mri_lesion_seg_app/rrunet3D.py. |
| Development-validation metrics | Approval Form for Artificial Intelligence Product – filled.docx, Section 5, item 4; same values reproduced in SIG draft Table 2. |
| Cohort and locked-test sizes | SIG draft Table 1 and AI Approval Form cohort table: 264 whole-gland and 209 lesion locked-test cases. |
| Locked-test result status | No evaluation/organ or evaluation/lesion summary artefact located; HIVE notebooks and README state locked testing should occur once after model choices are frozen. |
| Urofusion compatibility | docs/UROFUSION_RTSTRUCT_ACCEPTANCE_2026-09-27.md and associated RTSTRUCT validation record. |
| CARPL status | SIG draft Problems 4–6 and milestones; current integration material is preparatory/local rather than evidence of completed live CARPL deployment. |