# Prostate MRI Inference Pipeline — HIVE Batch Output Report

Report Date: 22 July 2026

Environment: HIVE Linux PC (SGH)

Prepared for: Claude AI Agent (Main Windows PC)

# 1. Executive Summary

This report summarises the batch inference results from running the Prostate MRI Lesion Segmentation and Classification pipeline (MONAI Deploy) on 3,538 SGH de-identified prostate MRI cases on the HIVE Linux PC. The pipeline was deployed via Docker container with GPU acceleration. This document is intended to brief the Claude AI agent on the Main Windows PC about the full context, issues discovered, and tasks that need to be done for post-processing, re-training, and pipeline improvements.

# 2. Pipeline Overview

The pipeline takes multi-parametric prostate MRI DICOM studies (T2-weighted, ADC, and High b-value) as input and produces:

NIfTI files: t2.nii.gz, adc.nii.gz, highb.nii.gz (intermediate inputs)

Organ segmentation: organ.nii.gz (multi-class: 0=background, 1=TZ, 2=PZ)

Lesion segmentation: 5 per-fold probability maps, merged probability, final binary mask

PI-RADS classification: lesions.txt with per-lesion PI-RADS scores (2-5), volumes, major axis lengths

RTSTRUCT files: organ_RTSTRUCT.dcm, lesion_RTSTRUCT.dcm, combined_organ_lesion_RTSTRUCT.dcm

DICOM copies of pipeline-selected T2/ADC/HIGHB series for viewer import

Key pipeline characteristics:

Organ segmentation: Sliding-window inference at 1.0mm isotropic with TorchScript model

Lesion segmentation: 5-fold 3D RR-UNet ensemble at 0.5mm isotropic, organ-masked, threshold=0.63

Lesion classification: 3D ResNet on 64×64×64 crops, 4 classes → PI-RADS 2-5

High b-value filtering: Automatic extraction of highest b-value from multi-b DWI series

Series selection: Regex-based rules validated at 100% precision/recall on 43,470 annotated rows

# 3. Batch Results Summary (3,538 Cases)

| Status | Count | Percentage |
| --- | --- | --- |
| SUCCESS_COMPLETE | 3,475 | 98.2% |
| FAIL_RUNTIME | 62 | 1.8% |
| SUCCESS_EMPTY_LESION | 1 | 0.03% |
| Total | 3,538 | 100% |

The 62 FAIL_RUNTIME cases experienced container exit code 1 with no output produced — all were missing T2, ADC, and HIGHB modalities, indicating the DICOM series selector could not match any valid series. These are likely genuinely incompatible or malformed DICOM studies.

The 1 SUCCESS_EMPTY_LESION case had an empty organ mask (no prostate detected), so lesion inference was intentionally skipped and an all-zero lesion_mask.nii.gz was written.

# 4. Critical Issue: Multi-Island Organ & Lesion Segmentation

## 4.1 Problem Description

Dr Anh Tuan Doan has identified that many organ NIfTI segmentations and their resulting RTSTRUCTs contain multiple disconnected "islands" — small erroneous mask fragments that are NOT part of the prostate gland. These appear as scattered blobs separate from the main prostate mass. When converted to RTSTRUCT, each island becomes a separate contour, making the output cluttered and clinically misleading.

Dr Anh specifically flagged two cases (20240903_45fd973241d1ff6efe2e and 20240904_f2dd2744c5965cdfaf0e) and requested investigation and cleanup.

## 4.2 Scale of the Problem — Connected Component Analysis

A full connected-component analysis was performed on all 3,475 successful cases. The results reveal this is a systemic issue affecting nearly a third of all cases:

### Organ Segmentation Island Distribution:

| Island Count | Number of Cases | Notes |
| --- | --- | --- |
| 1 (clean) | 2,402 (69.1%) | Single connected prostate — ideal |
| 2 | 405 (11.7%) | One small fragment |
| 3-5 | 435 (12.5%) | Multiple small fragments |
| 6-10 | 183 (5.3%) | Moderate fragmentation |
| 11-20 | 42 (1.2%) | Severe fragmentation |
| 21-33 | 8 (0.2%) | Extreme fragmentation |
| Total multi-island | 1,073 (30.9%) | Need cleanup |

Of the 1,073 multi-island cases, 89 cases have SEVERE fragmentation where the largest connected component represents less than 90% of total organ voxels — meaning significant portions of the mask are erroneous islands.

### Lesion Segmentation Island Distribution:

While some cases legitimately have multiple lesion foci, extremely high island counts suggest false-positive fragmentary detections:

| Lesion Islands | Count |
| --- | --- |
| 0 (no lesion detected) | 706 |
| 1 | 1,682 |
| 2 | 744 |
| 3 | 228 |
| 4 | 75 |
| 5-15 | 40 |

# 5. Dr Anh's Specific Requests

## 5.1 Organ Mask Cleanup — Remove Erroneous Islands

Request: Implement an automatic cleanup algorithm to remove disconnected mask fragments that are not part of the prostate gland. Keep only the largest connected component.

Dr Anh suggested approach:

Use MONAI's KeepLargestConnectedComponent transform on the organ mask

Apply per-class (label=1 for TZ, label=2 for PZ) or binary (any non-zero) connected component filtering

The largest island IS the prostate; all smaller islands are erroneous

Apply BEFORE RTSTRUCT conversion to prevent fragmented contours

Consider also filling holes in the largest island for cleaner contours

Gemini AI script provided by Dr Anh (pseudocode):

1. Load NIfTI mask → 2. Find all connected "islands" of label-1 voxels → 3. Keep ONLY the largest island → 4. Create clean mask with just that island → 5. Calculate volume = (prostate_voxels × voxel_spacing_product) / 1000 cc → 6. Calculate bounding box dimensions in mm → 7. Save cleaned mask

## 5.2 Prostate Dimensions & Volume Calculation

Request: Calculate the 3 maximal orthogonal dimensions (AP, LR, SI in mm) and true volume (in mm³ and cc) from the cleaned/contoured prostate gland. The bounding-box approach from the NIfTI affine + voxel spacing is sufficient.

Dr Anh mentioned these could later be stored as DICOM SR (Structured Report) via CARPL, but the immediate need is to generate them from NIfTI.

## 5.3 Potential DICOM SR Integration

Future consideration: Encode the prostate dimensions and volume into a DICOM Structured Report (SR) for integration with CARPL viewer. This is lower priority than the cleanup and calculations.

# 6. Interesting Outputs Folder — 25 Selected Cases

A curated folder of 25 representative cases has been prepared for transfer to the Main Windows PC. Total size: ~1.6 GB. These cases are categorised into three groups for development, testing, and validation:

## 6.1 Dr Anh-Flagged Cases (2)

| Case ID | Issue | Notes |
| --- | --- | --- |
| 20240903_45fd973241d1ff6efe2e | Multi-island organ, fragmented RTSTRUCT | Prostate major axis 417mm (inflated by islands), 2 lesions detected PI-RADS 3 & 4 |
| 20240904_f2dd2744c5965cdfaf0e | Multi-island organ, fragmented RTSTRUCT | Prostate major axis 415mm (inflated by islands), no lesion detected |

## 6.2 Multi-Island Problem Cases (8)

Selected to represent the worst and most varied island fragmentation patterns, covering both organ-only and combined organ+lesion island problems:

| Case ID | Organ Islands | Lesion Islands | Key Feature |
| --- | --- | --- | --- |
| 20240204_fee2d51144a207fb2cdd | 33 | 0 | Most organ islands in entire cohort; largest=99.7% |
| 20231214_447fc7d5da37a4ecaf03 | 28 | 0 | 2nd-largest island is 14,127 voxels (14% of total) |
| 20231101_caa2e360b1aaa0f4c65d | 20 | 4 | WORST fragmentation: largest island only 41% of total |
| 20230724_9842be66f7f42676b51b | 17 | 2 | Largest organ island=76.4%; 2nd island=14,240 vox |
| 20250806_eb53e15da0a87869aa59 | 25 | 4 | Both organ AND lesion multi-island |
| 20250226_e49e263e4fee78681a84 | 15 | 6 | Both organ AND lesion multi-island |
| 20230221_272da2e6002c60fb7921 | 14 | 15 | WORST lesion fragmentation in entire cohort |
| 20241118_967385a01ba9a4bd3871 | 2 | 9 | Minimal organ issue but extreme lesion fragmentation |

## 6.3 Perfect Reference Cases (15)

These cases have single-island organ segmentation, detected lesion(s), all 3 RTSTRUCT files, and complete pipeline outputs. They represent what "correct" output looks like and can be used as positive reference for testing cleanup algorithms and validating pipeline changes:

| Case ID | Lesions Detected | PI-RADS | Notes |
| --- | --- | --- | --- |
| 20230221_823319ef014cb44958c1 | 1 | 4 | Small prostate (7,431 mm³) |
| 20230221_846949d736c10b2b1350 | 1 | 4 | Medium prostate (8,206 mm³) |
| 20230222_2f4df4a94ccd37425b48 | 1 | 5 | Large lesion (3,795 mm³) |
| 20230222_cb66cb41b9efbe8cd217 | 1 | 5 | Medium prostate, small lesion |
| 20230222_ecdb6be112472ce42c74 | 1 (2 components) | — | 2 lesion islands, single organ |
| 20230223_8f6c60bf074349e7f155 | 1 (2 components) | — | 2 lesion islands, single organ |
| 20230224_57f45ff07453250f8b8d | 1 | 5 | Medium-large prostate (10,408 mm³) |
| 20230224_d3f2e3fd403e7500936f | 1 | 3 | Small prostate, PI-RADS 3 only |
| 20230227_023354fdbf7a426bc2ad | 1 | 5 | Small prostate, large lesion-to-organ ratio |
| 20230228_38a45cafa0b65b82999b | 1 (2 components) | — | 2 lesion islands |
| 20230228_4d25cbcc3e587816c9e0 | 1 (2 components) | — | 2 lesion islands, medium prostate |
| 20230301_1439ac3c5b84973439e3 | 2 | 3 + 5 | Two distinct lesions, different PI-RADS |
| 20230301_b25646560f3dbd4ad126 | 2 | 4 + 4 | Two lesions, same PI-RADS grade |
| 20230302_8a07f53b06cba14bd7e9 | 1 | 3 | PI-RADS 3 only case |
| 20230307_4d34fbad0f077c83407d | 1 | 3 | PI-RADS 3 only case |

# 7. Tasks for Claude AI Agent on Main Windows PC

The following work needs to be done on the Main Windows PC using these 25 interesting cases and the pipeline codebase. The HIVE Linux PC is for inference only — all development, re-training, and pipeline modifications happen on the Main Windows PC.

## 7.1 Priority 1 — Organ Mask Cleanup Post-Processing

Implement a post-processing step in the pipeline to automatically remove erroneous organ mask islands:

Add KeepLargestConnectedComponent (from MONAI) after organ segmentation inference, BEFORE RTSTRUCT generation

Apply independently to each organ label class (TZ=1, PZ=2) or binarise first, keep largest, then restore classes

Test on the 8 multi-island problem cases — verify islands are removed and only true prostate remains

Test on the 15 perfect cases — verify cleanup does NOT remove any valid prostate voxels

Consider adding morphological hole-filling (scipy.ndimage.binary_fill_holes) after island removal for smoother contours

Update organ_seg_operator.py or add a new cleanup operator in the DAG

Regenerate RTSTRUCTs from cleaned organ mask and verify they display correctly

## 7.2 Priority 2 — Lesion Mask Cleanup

Consider applying similar cleanup to lesion masks:

Remove very tiny lesion islands (e.g., <5 voxels) that are likely noise

Be careful NOT to remove legitimate small lesion foci — clinical significance threshold needed

Test on the multi-lesion problem cases (especially 20230221_272da2e6002c60fb7921 with 15 lesion islands)

Consider minimum-volume threshold rather than keep-largest-only (since patients can have multiple real lesions)

## 7.3 Priority 3 — Prostate Volume & Dimensions Calculation

Add prostate gland measurements to the output:

Calculate 3 maximal orthogonal dimensions (AP, LR, SI) in mm from the cleaned organ mask

Calculate true volume in mm³ and cc using voxel spacing from NIfTI affine

Add these measurements to lesions.txt (YAML report) alongside existing organ stats

Currently, the Major_Axis_Length in lesions.txt appears inflated for island cases (e.g., 417mm when actual prostate should be ~50-80mm) — this MUST be recalculated from cleaned mask

Consider generating DICOM SR (Structured Report) for CARPL integration (lower priority)

## 7.4 Priority 4 — Pipeline Integration & Re-training Considerations

Integrate cleanup into the MONAI Deploy DAG so it runs automatically for all future inferences

Consider whether the organ model itself needs re-training with cleaner training data to reduce islands at the source

If re-training, the 1,073 multi-island cases can provide training signal — clean up their organ masks and use as additional training data

The 2,402 clean single-island cases can be used directly as-is for training

Validate that cleanup does not adversely affect downstream lesion segmentation (since lesion ROI is derived from organ mask)

Update batch_audit_log.csv schema to include island counts and cleanup status for future batch runs

# 8. Output Folder Structure Reference

Each case folder in the Interesting outputs directory contains:

adc/adc.nii.gz — Intermediate NIfTI (ADC map)

t2/t2.nii.gz — Intermediate NIfTI (T2-weighted)

highb/highb.nii.gz — Intermediate NIfTI (High b-value)

organ/organ.nii.gz — Multi-class prostate mask (0=bg, 1=TZ, 2=PZ)

organ/organ_RTSTRUCT.dcm — Organ RTSTRUCT

lesion/lesion_mask.nii.gz — Final binary lesion mask

lesion/merged_lesion_prob.nii.gz — Averaged probability map

lesion/fold{0-4}_lesion_prob.nii.gz — Per-fold probability maps

lesion/lesion_RTSTRUCT.dcm — Lesion RTSTRUCT

combined_organ_lesion_RTSTRUCT.dcm — Combined organ+lesion RTSTRUCT

lesions.txt — PI-RADS report (YAML format)

dicom/{t2,adc,highb}/ — Copies of pipeline-selected DICOM series

# 9. Technical Context for the Claude AI Agent

## 9.1 Key Code Files to Modify

| File | Purpose |
| --- | --- |
| prostate_mri_lesion_seg_app/organ_seg_operator.py | Organ segmentation — add post-processing cleanup here |
| prostate_mri_lesion_seg_app/custom_lesion_seg_operator.py | Lesion segmentation — consider tiny-island removal |
| prostate_mri_lesion_seg_app/custom_lesion_classifier_operator.py | Classifier — add volume/dimension calculations |
| prostate_mri_lesion_seg_app/rtstruct_utils.py | RTSTRUCT generation — will automatically benefit from cleaned masks |
| prostate_mri_lesion_seg_app/app.py | Main DAG — may need new operator for cleanup step |

## 9.2 Key MONAI Transforms to Use

KeepLargestConnectedComponent(applied_labels=[1], is_onehot=False) — remove small islands

FillHoles() — fill holes in the largest component for smoother contours

generate_spatial_bounding_box() — get bounding box for dimension calculation

SaveImage() — save cleaned mask back to NIfTI

## 9.3 Important Notes

The organ mask is MULTI-CLASS (0=bg, 1=TZ, 2=PZ). Cleanup must handle both classes correctly — either apply per-class or binarise→cleanup→restore classes.

Lesion inference is organ-masked — any changes to the organ mask will affect the lesion ROI.

The current Major_Axis_Length in lesions.txt is calculated from the RAW organ mask including islands. After cleanup, this value will be corrected automatically.

RTSTRUCT generation uses rt_utils and references the T2 DICOM series by SeriesInstanceUID. No changes needed to RTSTRUCT code — just feed it cleaned NIfTI masks.

The pipeline threshold for lesion is fixed at 0.63. This was not part of Dr Anh's requests but could be revisited if lesion island analysis shows too many false positives.

All 3,538 cases' audit log (batch_audit_log.csv) is in the output/ folder on the HIVE PC for reference.

# 10. Appendix: Full Statistical Summary

## 10.1 Case Outcomes

Total DICOM study folders processed: 3,538 (from SGH de-identified data)

SUCCESS_COMPLETE: 3,475 (98.2%) — all modalities found, organ + lesion inference ran, classifier report produced

FAIL_RUNTIME: 62 (1.8%) — container exited code 1, no output (all missing T2/ADC/HIGHB)

SUCCESS_EMPTY_LESION: 1 (0.03%) — organ mask empty, lesion skipped

## 10.2 Lesion Detection

Cases with at least one lesion detected: 2,769 out of 3,475 (79.7%)

Cases with no lesion detected: 706 (20.3%) — organ present but no lesion above threshold

## 10.3 Organ Segmentation Quality

Single-island (clean) organ: 2,402 (69.1%) — ideal output

Multi-island organ: 1,073 (30.9%) — needs KeepLargestConnectedComponent cleanup

Severe fragmentation (largest <90%): 89 (2.6%) — most urgent cleanup candidates

## 10.4 Interesting Outputs Manifest

The Interesting outputs folder also contains a MANIFEST.json file with category tags for each case (DR_ANH_FLAGGED, MULTI_ISLAND_PROBLEM, or PERFECT_REFERENCE).